## Slide 1: AAPM Computed Tomography Automatic Exposure Control Education Slides
Many of the terms used in these slides can be found in the CT Terminology Lexicon
http://www.aapm.org/pubs/CTProtocols/documents/CTTerminologyLexicon.pdf
Last updated 01/15/2021

## Slide 2: Disclaimer
Screen captures are examples of a common (or latest) software version only and all software versions are not represented
The information contained herein is current as of the date shown on the title slide
Modification of the content of these slides is NOT allowed.
  The modified content, including indirect or unintentional changes in the accuracy or meaning of related content, becomes the sole responsibility of the person/organization creating and/or using the edited version.
  Neither the AAPM nor the manufacturers participating in creating this slide set assume any responsibility for edited versions of these slides, or for content of oral presentations associated with the original or edited slides.

## Slide 3: Motivation
These slides are provided to aid in understanding the factors that affect performance of Automatic Exposure Control, specifically image quality and radiation dose, in CT studies

## Slide 4: Outline
Effect of CT localizer on AEC
Image quality reference parameter for AEC
Effect of patient size on AEC
Effect of scanned anatomy
Effect of first or expected reconstruction settings
Advanced AEC features

## Slide 5: Effect of CT Localizer
The CT localizer(s) provide the initial data to inform the behavior of the AEC
The apparent size of the patient on the localizer(s) or the measured attenuation are used to set the initial dose level for the exam
The localizer(s) may also be used to adjust the longitudinal or angular tube current modulation
The use of multiple localizers and the order of their acquisition may affect the behavior of the system’s AEC

## Slide 6: Image quality reference parameter for AEC
The image quality reference parameter for AEC is generally a measure of image quality in the reconstructed images
The image quality reference parameter for AEC has a unique relationship with both tube output and patient size
Specifically, the Image quality reference parameter is used together with the patient attenuation profile (as estimated by the CT localizer) to determine the tube output for a particular exam
The operation of the AEC may be independent of the reconstruction parameters, or related to them

## Slide 7: Advanced AEC Features Outline
AEC in cardiac exams
Unusual attenuation profiles
  Head/Neck exams (strategy to handle abrupt change of attenuation profile)
  Extremity exams
  Neonates and very small children
  Metal/Foreign objects within Scan FOV
  Obese patients
Automatic tube voltage selection
Organ based tube current modulation